## LabView VI
Connecting Chassis | Connector Type on Slow Interface Panel | Signals | Pin | Signal Type | Terminated to | NI PXI-6229 Pins | Bypass Connector
PXI-6229 Unit 1
STS-2 Interface - 1 | D-sub, 15-pin, M | U | +1, -9 | analog in | internal slow control | ai-0 |  | A1:PEM-SEI_STS2_1_U_VOLT
|  | V | +2, -10 | analog in | internal slow control | ai-1 |  | A1:PEM-SEI_STS2_1_V_VOLT
|  | W | +3, -11 | analog in | internal slow control | ai-2 |  | A1:PEM-SEI_STS2_1_W_VOLT
|  | X | +4, -12 | analog in | pass thru to AA |  | D-sub, 9-pin, F, Connectro 1 - pin 1,6
|  | Y | +5, -13 | analog in | pass thru to AA |  | D-sub, 9-pin, F, Connectro 1 - pin 2,7
|  | Z | +6, -14 | analog in | pass thru to AA |  | D-sub, 9-pin, F, Connectro 1 - pin 3,8
| D-sub, 9-pin, F | CalSW |  | digital out | -
|  | SigSel |  | digital out | -
|  | Period |  | digital out | -
|  | Autozero |  | digital out | internal slow control | p1-0 |  | A1:PEM-SEI_STS2_1_AUTOZERO
STS-2 Interface - 2 | D-sub, 15-pin, M | U |  | analog in | internal slow control | ai-3 |  | A1:PEM-SEI_STS2_2_U_VOLT
|  | V |  | analog in | internal slow control | ai-4 |  | A1:PEM-SEI_STS2_2_V_VOLT
|  | W |  | analog in | internal slow control | ai-5 |  | A1:PEM-SEI_STS2_2_W_VOLT
|  | X |  | analog in | pass thru to AA |  | D-sub, 9-pin, F, Connectro 1 - pin 4,9
|  | Y |  | analog in | pass thru to AA |  | D-sub, 9-pin, F, Connectro 2 - pin 1,6
|  | Z |  | analog in | pass thru to AA |  | D-sub, 9-pin, F, Connectro 2 - pin 2,7
| D-sub, 9-pin, F | CalSW |  | digital out | -
|  | SigSel |  | digital out | -
|  | Period |  | digital out | -
|  | Autozero |  | digital out | internal slow control | p1-1 |  | A1:PEM-SEI_STS2_2_AUTOZERO
ISC VGA Chassis | D-sub, 37-pin, F | 24dB gain, Ch 1 |  | digital out | internal slow control | p0-0 |  | A1:PXI-LSC_VGA_CH1_GAIN_DB
|  | 12dB gain, Ch 1 |  | digital out | internal slow control | p0-1
|  | 6dB gain, Ch 1 |  | digital out | internal slow control | p0-2
|  | 3dB gain, Ch 1 |  | digital out | internal slow control | p0-3
|  | PoleZero1, Ch 1 |  | digital out | internal slow control | p0-4 |  | A1:PXI-LSC_VGA_CH1_PZ1
|  | PoleZero2, Ch 1 |  | digital out | internal slow control | p0-5
|  | PoleZero3, Ch 1 |  | digital out | internal slow control | p0-6
|  | LE, ch1 |  | digital out | internal slow control | p0-7
|  | 24dB gain, Ch 2 |  | digital out | internal slow control | p0-8
|  | 12dB gain, Ch 2 |  | digital out | internal slow control | p0-9
|  | 6dB gain, Ch 2 |  | digital out | internal slow control | p0-10
|  | 3dB gain, Ch 2 |  | digital out | internal slow control | p0-11
|  | PoleZero1, Ch 2 |  | digital out | internal slow control | p0-12
|  | PoleZero2, Ch 2 |  | digital out | internal slow control | p0-13
|  | PoleZero3, Ch 2 |  | digital out | internal slow control | p0-14
|  | LE, ch2 |  | digital out | internal slow control | p0-15
|  | 24dB gain, Ch 3 |  | digital out | internal slow control | p0-16
|  | 12dB gain, Ch 3 |  | digital out | internal slow control | p0-17
|  | 6dB gain, Ch 3 |  | digital out | internal slow control | p0-18
|  | 3dB gain, Ch 3 |  | digital out | internal slow control | p0-19
|  | PoleZero1, Ch 3 |  | digital out | internal slow control | p0-20
|  | PoleZero2, Ch 3 |  | digital out | internal slow control | p0-21
|  | PoleZero3, Ch 3 |  | digital out | internal slow control | p0-22
|  | LE, ch 3 |  | digital out | internal slow control | p0-23
|  | 24dB gain, Ch 4 |  | digital out | internal slow control | p0-24
|  | 12dB gain, Ch 4 |  | digital out | internal slow control | p0-25
|  | 6dB gain, Ch 4 |  | digital out | internal slow control | p0-26
|  | 3dB gain, Ch 4 |  | digital out | internal slow control | p0-27
|  | PoleZero1, Ch 4 |  | digital out | internal slow control | p0-28
|  | PoleZero2, Ch 4 |  | digital out | internal slow control | p0-29
|  | PoleZero3, Ch 4 |  | digital out | internal slow control | p0-30
|  | LE, ch 4 |  | digital out | internal slow control | p0-31
TTFSS - 1 | D-sub, 25-pin, MF | Comm Gain In |  | analog out | internal slow control | ao-0 |  | A1:PXI-PSL_FSS_LASER_A_COMM_GAIN
|  | Offset In |  | analog out | internal slow control | ao-1 |  | A1:PXI-PSL_FSS_LASER_A_OFFSET
|  | Fast Gain In |  | analog out | internal slow control | ao-2 |  | A1:PXI-PSL_FSS_LASER_A_FAST_GAIN
|  | Slow In |  | analog out | internal slow control | ao-3 |  | A1:PXI-PSL_FSS_LASER_A_SLOW_GAIN
|  | Slow Out |  | analog in |  | ai-6 |  | A1:PXI-PSL_FSS_LASER_A_SLOW_OUT
|  | LO mon |  | analog in |  | ai-7 |  | A1:PXI-PSL_FSS_LASER_A_LO_OUT
|  | Mixer mon |  | analog in |  | ai-8 |  | A1:PXI-PSL_FSS_LASER_A_MIXER_OUT
|  | EOM mon |  | analog in |  | ai-9 |  | A1:PXI-PSL_FSS_LASER_A_EOM_OUT
|  | Fast mon |  | analog in |  | ai-10 |  | A1:PXI-PSL_FSS_LASER_A_FAST_OUT
|  | Slow Enable |  | digital out |  | p1-2 |  | A1:PXI-PSL_FSS_LASER_A_SLOW_EN
|  | T1 Enable |  | digital out |  | p1-3 |  | A1:PXI-PSL_FSS_LASER_A_TEST1_EN
|  | Ramp Enable |  | digital out |  | p1-4 |  | A1:PXI-PSL_FSS_LASER_A_RAMP_EN
|  | T2 Enable |  | digital out |  | p1-5 |  | A1:PXI-PSL_FSS_LASER_A_TEST2_EN
|  | Additive Offset Enable |  | digital out |  | p1-6 |  | A1:PXI-PSL_FSS_LASER_A_AO_EN
|  | Remote mon |  | digital in |  | p1-7
ai-11
ai-12
ai-13
ai-14
ai-15
PXI6229 - Unit 2
STS-2 Interface - 3 | D-sub, 15-pin, M | U |  | analog in | internal slow control | ai-0 |  | A1:PEM-SEI_STS2_3_U_VOLT
|  | V |  | analog in | internal slow control | ai-1 |  | A1:PEM-SEI_STS2_3_V_VOLT
|  | W |  | analog in | internal slow control | ai-2 |  | A1:PEM-SEI_STS2_3_W_VOLT
|  | X |  | analog in | pass thru to AA |  | D-sub, 9-pin, MF, Connectro 2 - pin 3,8
|  | Y |  | analog in | pass thru to AA |  | D-sub, 9-pin, MF, Connectro 2 - pin 4,9
|  | Z |  | analog in | pass thru to AA |  | D-sub, 9-pin, MF, Connectro 3 - pin 1,6
| D-sub, 9-pin, F | CalSW |  | digital out | -
|  | SigSel |  | digital out | -
|  | Period |  | digital out | -
|  | Autozero |  | digital out | internal slow control | p1-0 |  | A1:PEM-SEI_STS2_3_AUTOZERO
STS-2 Interface - 4 | D-sub, 15-pin, M | U |  | analog in | internal slow control | ai-3 |  | A1:PEM-SEI_STS2_4_U_VOLT
|  | V |  | analog in | internal slow control | ai-4 |  | A1:PEM-SEI_STS2_4_V_VOLT
|  | W |  | analog in | internal slow control | ai-5 |  | A1:PEM-SEI_STS2_4_W_VOLT
|  | X |  | analog in | pass thru to AA |  | D-sub, 9-pin, MF, Connectro 3 - pin 2,7
|  | Y |  | analog in | pass thru to AA |  | D-sub, 9-pin, MF, Connectro 3 - pin 3,8
|  | Z |  | analog in | pass thru to AA |  | D-sub, 9-pin, MF, Connectro 3 - pin 4,9
| D-sub, 9-pin, F | CalSW |  | digital out | -
|  | SigSel |  | digital out | -
|  | Period |  | digital out | -
|  | Autozero |  | digital out | internal slow control | p1-1 |  | A1:PEM-SEI_STS2_4_AUTOZERO
TTFSS - 2 | D-sub, 25-pin, MF | Comm Gain In |  | analog out | internal slow control | ao-0 |  | A1:PXI-PSL_FSS_LASER_B_COMM_GAIN
|  | Offset In |  | analog out | internal slow control | ao-1 |  | A1:PXI-PSL_FSS_LASER_B_OFFSET
|  | Fast Gain In |  | analog out | internal slow control | ao-2 |  | A1:PXI-PSL_FSS_LASER_B_FAST_GAIN
|  | Slow In |  | analog out | internal slow control | ao-3 |  | A1:PXI-PSL_FSS_LASER_B_SLOW_GAIN
|  | Slow Out |  | analog in |  | ai-6 |  | A1:PXI-PSL_FSS_LASER_B_SLOW_OUT
|  | LO mon |  | analog in |  | ai-7 |  | A1:PXI-PSL_FSS_LASER_B_LO_OUT
|  | Mixer mon |  | analog in |  | ai-8 |  | A1:PXI-PSL_FSS_LASER_B_MIXER_OUT
|  | EOM mon |  | analog in |  | ai-9 |  | A1:PXI-PSL_FSS_LASER_B_EOM_OUT
|  | Fast mon |  | analog in |  | ai-10 |  | A1:PXI-PSL_FSS_LASER_B_FAST_OUT
|  | Slow Enable |  | digital out |  | p1-2 |  | A1:PXI-PSL_FSS_LASER_B_SLOW_EN
|  | T1 Enable |  | digital out |  | p1-3 |  | A1:PXI-PSL_FSS_LASER_B_TEST1_EN
|  | Ramp Enable |  | digital out |  | p1-4 |  | A1:PXI-PSL_FSS_LASER_B_RAMP_EN
|  | T2 Enable |  | digital out |  | p1-5 |  | A1:PXI-PSL_FSS_LASER_B_TEST2_EN
|  | Additive Offset Enable |  | digital out |  | p1-6 |  | A1:PXI-PSL_FSS_LASER_B_AO_EN
|  | Remote mon |  | digital in |  | p1-7
Temperature Sensor - 1 | D-sub, 9-pin, F | RTS |  | analog in |  | ai-11
|  | RTS |  | analog in |  | ai-12
|  | RTS |  | analog in |  | ai-13
|  | +5V |  | +5V out
Temperature Sensor - 2 | D-sub, 9-pin, F | RTS |  | analog in |  | ai-14
|  | RTS |  | analog in |  | ai-15
|  | RTS |  | analog in
|  | +5V |  | +5V out

## Sheet2
Gain |  |  |  |  | 37-pin D-Sub Digital Input - VGA
0 |  |  |  |  | 37 pin sub-D | ribbon cable wire color | VGA Signals | VGA Signals
3 |  |  |  |  | 1 | black | GND
6 |  |  |  |  | 20 | white |  | GND
9 |  |  |  |  | 2 | grey | BI-0
12 |  |  |  |  | 21 | purple |  | BI-16
15 |  |  |  |  | 3 | blue | BI-1
18 |  |  |  |  | 22 | green |  | BI-17
21 |  |  |  |  | 4 | yellow | BI-2
24 |  |  |  |  | 23 | orange |  | BI-18
27 |  |  |  |  | 5 | red | BI-3
30 |  |  |  |  | 24 | brown |  | BI-19
33 |  |  |  |  | 6 | black | BI-4
36 |  |  |  |  | 25 | white |  | BI-20
39 |  |  |  |  | 7 | grey | BI-5
42 |  |  |  |  | 26 | purple |  | BI-21
45 |  |  |  |  | 8 | blue | BI-6
[5 empty] | 27 | green |  | BI-22
gain | 3 |  |  |  | 9 | yellow | BI-7
[5 empty] | 28 | orange |  | BI-23
|  |  | bit pattern | decimal | 10 | red | BI-8
bit 0 | 24 | 0 | 0 | 0 | 29 | brown |  | BI-24
bit 1 | 12 | 0 | 0 | 0 | 11 | black | BI-9
bit 2 | 6 | 0 | 0 | 0 | 30 | white |  | BI-25
bit 3 | 3 | 0 | 0 | 0 | 12 | grey | BI-10
bit 4 | 0 | 0 | 0 | 0 | 31 | purple |  | BI-26
bit 5 | 0 | 0 | 0 | 0 | 13 | blue | BI-11
bit 6 | 0 | 0 | 0 | 0 | 32 | green |  | BI-27
bit 7 | 1 | 1 | 1 | 128 | 14 | yellow | BI-12
[5 empty] | 33 | orange |  | BI-28
|  |  |  | 128 | 15 | red | BI-13
[5 empty] | 34 | brown |  | BI-29
[5 empty] | 16 | black | BI-14
[5 empty] | 35 | white |  | BI-30
[5 empty] | 17 | grey | BI-15
[5 empty] | 36 | purple |  | BI-31
[5 empty] | 18 | blue | +15V
[5 empty] | 37 | green |  | +15V
[5 empty] | 19 | yellow | NC

## ACROMAG Units
ADC Channels | 39 | ADC-x | - | ch# | EPICS Channel Name | ACROMAT Unit Pins | Front Panel Connect | ACROMAG |  |  |  |  | QTY
STS-2-A | 3 | ADC0 | Read | 0 | A1:PEM-STS2_A_MASS_POS_U |  |  | XT1111 | 16-channel IO |  |  |  | 4
|  |  | Read | 1 | A1:PEM-STS2_A_MASS_POS_V |  |  | XT1221 | 8-channel diff AI |  |  |  | 5
|  |  | Read | 2 | A1:PEM-STS2_A_MASS_POS_W |  |  | XT1541 | 8-channel AO |  |  |  | 2
STS-2-B | 3 |  | Read | 3 | A1:PEM-STS2_B_MASS_POS_U |  |  | 966EN-4006 | 6-channel RTD/Resistance Input |  |  |  | 1
|  |  | Read | 4 | A1:PEM-STS2_B_MASS_POS_V
|  |  | Read | 5 | A1:PEM-STS2_B_MASS_POS_W
STS2-C | 3 |  | Read | 6 | A1:PEM-STS2_C_MASS_POS_U
|  |  | Read | 7 | A1:PEM-STS2_C_MASS_POS_V |  |  | SEI Slow |  |  |  |  | QTY
|  | ADC1 | Read | 0 | A1:PEM-STS2_C_MASS_POS_W |  |  | XT1111 | 16-channel IO |  |  |  | 3
STS-2-D | 3 |  | Read | 1 | A1:PEM-STS2_D_MASS_POS_U |  |  | XT1221 | 8-channel diff AI |  |  |  | 3
|  |  | Read | 2 | A1:PEM-STS2_D_MASS_POS_V |  |  | 966EN-4006 | 6-channel RTD/Resistance Input |  |  |  | 1
|  |  | Read | 3 | A1:PEM-STS2_D_MASS_POS_W
T240-A | 4 |  | Read | 4 | A1:ISI-T240_A_MASS_POS_U
|  |  | Read | 5 | A1:ISI-T240_A_MASS_POS_V
|  |  | Read | 6 | A1:ISI-T240_A_MASS_POS_W
|  |  | Read | 7 | A1:ISI-T240_A_POD_PRESSURE_MBAR |  |  | ISC Slow |  |  | Req # | Spare # | Qty
T240-B | 4 | ADC2 | Read | 0 | A1:ISI-T240_B_MASS_POS_U |  |  | XT1111 | 16-channel IO |  | 12 | 4 | 1
|  |  | Read | 1 | A1:ISI-T240_B_MASS_POS_V |  |  | XT1221 | 8-channel diff AI |  | 10 | 6 | 2
|  |  | Read | 2 | A1:ISI-T240_B_MASS_POS_W |  |  | XT1541 | 8-channel AO |  | 12 | 4 | 2
|  |  | Read | 3 | A1:ISI-T240_B_POD_PRESSURE_MBAR
T240-C | 4 |  | Read | 4 | A1:ISI-T240_C_MASS_POS_U
|  |  | Read | 5 | A1:ISI-T240_C_MASS_POS_V
|  |  | Read | 6 | A1:ISI-T240_C_MASS_POS_W |  |  | Hardware
|  |  | Read | 7 | A1:ISI-T240_C_POD_PRESSURE_MBAR |  |  | ATF Wiki Hardware listing |  | https://nodus.ligo.caltech.edu:30889/ATFWiki/doku.php?id=main:resources:computing:front_panel
TTSFF-A | 5 | ADC3 | Read | 0 | A1:ISC-FSS_LASER_A_SLOW_RB |  |  | 4U Chassis | http://au.element14.com/schroff/20860-612/rack-multipac-pro-4u-340-d-plain/dp/1816047
|  |  | Read | 1 | A1:ISC-FSS_LASER_A_LOM_RB |  |  | subD Terminals |  | http://www.ebay.com/itm/Slim-Right-Angle-D-039-SUB-Header-Breakout-Board-Terminal-Block-DSUB-Connector-/201456303236
|  |  | Read | 2 | A1:ISC-FSS_LASER_A_MIXR_RB |  |  | BNC |  | http://au.element14.com/amphenol-rf/031-10-rfxg1/connector-bnc-bhd-jack-str-50/dp/2396249
|  |  | Read | 3 | A1:ISC-FSS_LASER_A_PCM_RB |  |  | Ethernet feedthrough |  | http://au.element14.com/l-com/ecf504-sc5e/adaptor-rj45-jack-rj45-jack-coupler/dp/1704447?ost=21K6038&searchView=table&iscrfnonsku=false&ddkey=http%3Aen-AU%2FElement14_Australia%2Fsearch
|  |  | Read | 4 | A1:ISC-FSS_LASER_A_FAST_RB |  |  |  |  | http://www.l-com.com/ethernet-cat5e-rj45-coupler-shielded-8x8-panel-mount-style
TTSFF-B | 5 |  | Read | 5 | A1:ISC-FSS_LASER_B_SLOW_RB |  |  | Cat5E patch cable |  | http://au.element14.com/videk/2965-0-3b/patch-lead-cat5e-blue-0-3m/dp/1526152
|  |  | Read | 6 | A1:ISC-FSS_LASER_B_LOM_RB
|  |  | Read | 7 | A1:ISC-FSS_LASER_B_MIXR_RB
|  | ADC4 | Read | 0 | A1:ISC-FSS_LASER_B_PCM_RB
|  |  | Read | 1 | A1:ISC-FSS_LASER_B_FAST_RB
TTSFF-C | 5 |  | Read | 2 | A1:ISC-FSS_LASER_C_SLOW_RB
|  |  | Read | 3 | A1:ISC-FSS_LASER_C_LOM_RB
|  |  | Read | 4 | A1:ISC-FSS_LASER_C_MIXR_RB
|  |  | Read | 5 | A1:ISC-FSS_LASER_C_PCM_RB
|  |  | Read | 6 | A1:ISC-FSS_LASER_C_FAST_RB
|  |  | Read | 7 | SPARE
DAC Channels | 12 | DAC-x |  | Ch# | EPICS Channel name
TTSFF-A | 4 | DAC0 | WRITE | 0 | A1:ISC-FSS_LASER_A_COMM_GAIN_V
|  |  | WRITE | 1 | A1:ISC-FSS_LASER_A_OFFSET_V
|  |  | WRITE | 2 | A1:ISC-FSS_LASER_A_FAST_GAIN_V
|  |  | WRITE | 3 | A1:ISC-FSS_LASER_A_SLOW_GAIN_V
TTSFF-B | 4 |  | WRITE | 4 | A1:ISC-FSS_LASER_B_COMM_GAIN_v
|  |  | WRITE | 5 | A1:ISC-FSS_LASER_B_OFFSET_V
|  |  | WRITE | 6 | A1:ISC-FSS_LASER_B_FAST_GAIN_V
|  |  | WRITE | 7 | A1:ISC-FSS_LASER_B_SLOW_GAIN_V
TTSFF-C | 4 | DAC1 | WRITE | 0 | A1:ISC-FSS_LASER_C_COMM_GAIN_V
|  |  | WRITE | 1 | A1:ISC-FSS_LASER_C_OFFSET_V
|  |  | WRITE | 2 | A1:ISC-FSS_LASER_C_FAST_GAIN_V
|  |  | WRITE | 3 | A1:ISC-FSS_LASER_C_SLOW_GAIN_V
LASER-A | 1 |  | WRITE | 4 | A1:ISC-LASER_A_CRYSTAL_TEMP_VOLT
LASER-B | 1 |  | WRITE | 5 | A1:ISC-LASER_B_CRYSTAL_TEMP_VOLT
LASER-C | 1 |  | WRITE | 6 | A1:ISC-LASER_C_CRYSTAL_TEMP_VOLT
LASER-D | 1 |  | WRITE | 7 | A1:ISC-LASER_D_CRYSTAL_TEMP_VOLT
Digital Out | 43 | BO-x |  | Ch# | EPICs Channel Name
STS-2-A | 3 | BIO0 | READ | 0 | A1:ISI-T240_C_WCAL_ENABLE_RB
|  |  | WRITE | 1 | A1:PEM-STS2_A_SIGSEL |  | Unit 1 - Digital From STS2 - pin 2
|  |  | WRITE | 2 | A1:PEM-STS2_A_PERIOD |  | Unit 1 - Digital From STS2 - pin 3
|  |  | WRITE | 3 | A1:PEM-STS2_A_AUTOZERO |  | Unit 1 - Digital From STS2 - pin 4
STS-2-B | 4 |  | READ | 4 | A1:ISI-T240_A_GAIN_RB |  | Unit 2 - Digital From STS2 - pin 1, A1:PEM-STS2_B_CALSW
|  |  | WRITE | 5 | A1:PEM-STS2_B_SIGSEL |  | Unit 2 - Digital From STS2 - pin 2
|  |  | WRITE | 6 | A1:PEM-STS2_B_PERIOD |  | Unit 2 - Digital From STS2 - pin 3
|  |  | WRITE | 7 | A1:PEM-STS2_B_AUTOZERO |  | Unit 2 - Digital From STS2 - pin 4
STS-2-C | 4 |  | READ | 8 | A1:ISI-T240_B_GAIN_RB |  | Unit 3 - Digital From STS2 - pin 1, A1:PEM-STS2_C_CALSW
|  |  | WRITE | 9 | A1:PEM-STS2_C_SIGSEL |  | Unit 3 - Digital From STS2 - pin 2
|  |  | WRITE | 10 | A1:PEM-STS2_C_PERIOD |  | Unit 3 - Digital From STS2 - pin 3
|  |  | WRITE | 11 | A1:PEM-STS2_C_AUTOZERO |  | Unit 3 - Digital From STS2 - pin 4
STS-2-D | 4 |  | READ | 12 | A1:ISI-T240_C_GAIN_RB |  | Unit 4 - Digital From STS2 - pin 1, A1:PEM-STS2_D_CALSW
|  |  | WRITE | 13 | A1:PEM-STS2_D_SIGSEL |  | Unit 4 - Digital From STS2 - pin 2
|  |  | WRITE | 14 | A1:PEM-STS2_D_PERIOD |  | Unit 4 - Digital From STS2 - pin 3
|  |  | WRITE | 15 | A1:PEM-STS2_D_AUTOZERO |  | Unit 4 - Digital From STS2 - pin 4
T240-A | 6 | BIO1 | WRITE | 0 | A1:ISI-T240_A_SIGSEL
|  |  | WRITE | 1 | A1:ISI-T240_A_AUTOZERO
|  |  | WRITE | 2 | A1:ISI-T240_A_UCAL_ENABLE
|  |  | WRITE | 3 | A1:ISI-T240_A_VCAL_ENABLE
|  |  | WRITE | 4 | A1:ISI-T240_A_WCAL_ENABLE
|  |  | WRITE | 5 | A1:ISI-T240_A_GAIN
T240-B | 6 |  | WRITE | 6 | A1:ISI-T240_B_SIGSEL
|  |  | WRITE | 7 | A1:ISI-T240_B_AUTOZERO
|  |  | WRITE | 8 | A1:ISI-T240_B_UCAL_ENABLE
|  |  | WRITE | 9 | A1:ISI-T240_B_VCAL_ENABLE
|  |  | WRITE | 10 | A1:ISI-T240_B_WCAL_ENABLE
|  |  | WRITE | 11 | A1:ISI-T240_B_GAIN
T240-C | 6 |  | WRITE | 12 | A1:ISI-T240_C_SIGSEL
|  |  | WRITE | 13 | A1:ISI-T240_C_AUTOZERO
|  |  | WRITE | 14 | A1:ISI-T240_C_UCAL_ENABLE
|  |  | WRITE | 15 | A1:ISI-T240_C_VCAL_ENABLE
|  | BIO2 | WRITE | 0 | A1:ISI-T240_C_WCAL_ENABLE
|  |  | WRITE | 1 | A1:ISI-T240_C_GAIN
TTFSS-A | 5 | BIO3 | WRITE | 0 | A1:ISC-FSS_LASER_A_SLOW_EN
|  |  | WRITE | 1 | A1:ISC-FSS_LASER_A_TEST1_EN
|  |  | WRITE | 2 | A1:ISC-FSS_LASER_A_RAMP_EN
|  |  | WRITE | 3 | A1:ISC-FSS_LASER_A_TEST2_EN
|  |  | WRITE | 4 | A1:ISC-FSS_LASER_A_AO_EN
TTFSS-B | 5 |  | WRITE | 5 | A1:ISC-FSS_LASER_B_SLOW_EN
|  |  | WRITE | 6 | A1:ISC-FSS_LASER_B_TEST1_EN
|  |  | WRITE | 7 | A1:ISC-FSS_LASER_B_RAMP_EN
|  |  | WRITE | 8 | A1:ISC-FSS_LASER_B_TEST2_EN
|  |  | WRITE | 9 | A1:ISC-FSS_LASER_B_AO_EN
TTFSS-C | 5 |  | WRITE | 10 | A1:ISC-FSS_LASER_C_SLOW_EN
|  |  | WRITE | 11 | A1:ISC-FSS_LASER_C_TEST1_EN
|  |  | WRITE | 12 | A1:ISC-FSS_LASER_C_RAMP_EN
|  |  | WRITE | 13 | A1:ISC-FSS_LASER_C_TEST2_EN
|  |  | WRITE | 14 | A1:ISC-FSS_LASER_C_AO_EN
|  |  | IO | 15 | SPARE
VGA -Chassis | 32 | BIO4 | WRITE | 0-3 | A1:PXI-LSC_VGA_CH1_GAIN_DB
|  |  | WRITE | 4 | A1:PXI-LSC_VGA_CH1_PZ1
|  |  | WRITE | 5 | A1:PXI-LSC_VGA_CH1_PZ2
|  |  | WRITE | 6 | A1:PXI-LSC_VGA_CH1_PZ3
|  |  | WRITE | 7 | A1:PXI-LSC_VGA_CH1_LE
|  |  | WRITE | `8-11 | A1:PXI-LSC_VGA_CH_GAIN_DB
|  |  | WRITE | 12 | A1:PXI-LSC_VGA_CH2_PZ1
|  |  | WRITE | 13 | A1:PXI-LSC_VGA_CH2_PZ2
|  |  | WRITE | 14 | A1:PXI-LSC_VGA_CH2_PZ3
|  |  | WRITE | 15 | A1:PXI-LSC_VGA_CH2_LE
|  | BIO5 | WRITE | 0-3 | A1:PXI-LSC_VGA_CH3_GAIN_DB
|  |  | WRITE | 4 | A1:PXI-LSC_VGA_CH3_PZ1
|  |  | WRITE | 5 | A1:PXI-LSC_VGA_CH3_PZ2
|  |  | WRITE | 6 | A1:PXI-LSC_VGA_CH3_PZ3
|  |  | WRITE | 7 | A1:PXI-LSC_VGA_CH3_LE
|  |  | WRITE | `8-11 | A1:PXI-LSC_VGA_CH4_GAIN_DB
|  |  | WRITE | 12 | A1:PXI-LSC_VGA_CH4_PZ1
|  |  | WRITE | 13 | A1:PXI-LSC_VGA_CH4_PZ2
|  |  | WRITE | 14 | A1:PXI-LSC_VGA_CH4_PZ3
|  |  | WRITE | 15 | A1:PXI-LSC_VGA_CH4_LE
Digital In | 17 | BI-x |  | Ch# | EPICS Channel Name
T240-A | 5 | BIO2 | READ | 2 | A1:ISI-T240_A_SIGSEL_RB
|  |  | READ | 3 | A1:ISI-T240_A_AUTOZERO_RB
|  |  | READ | 4 | A1:ISI-T240_A_UCAL_ENABLE_RB
|  |  | READ | 5 | A1:ISI-T240_A_VCAL_ENABLE_RB
|  |  | READ | 6 | A1:ISI-T240_A_WCAL_ENABLE_RB
T240-B | 5 |  | READ | 7 | A1:ISI-T240_B_SIGSEL_RB
|  |  | READ | 8 | A1:ISI-T240_B_AUTOZERO_RB
|  |  | READ | 9 | A1:ISI-T240_B_UCAL_ENABLE_RB
|  |  | READ | 10 | A1:ISI-T240_B_VCAL_ENABLE_RB
|  |  | READ | 11 | A1:ISI-T240_B_WCAL_ENABLE_RB
T240-C | 5 |  | READ | 12 | A1:ISI-T240_C_SIGSEL_RB
|  |  | READ | 13 | A1:ISI-T240_C_AUTOZERO_RB
|  |  | READ | 14 | A1:ISI-T240_C_UCAL_ENABLE_RB
|  |  | READ | 15 | A1:ISI-T240_C_VCAL_ENABLE_RB
|  | BIO0 | READ | 0 | A1:ISI-T240_C_WCAL_ENABLE_RB
TTSFF-A | 1 | DAC0 | READ | 0 | A1:ISC-FSS_LASER_A_REMOTE_EN_RB
TTSFF-B | 1 | DAC0 | READ | 1 | A1:ISC-FSS_LASER_B_REMOTE_EN_RB
TTSFF-C | 1 | DAC0 | READ | 2 | A1:ISC-FSS_LASER_C_REMOTE_EN_RB
|  |  | IO | 3 | SPARE
|  | DAC1 | IO | 0 | SPARE
|  |  | IO | 1 | SPARE
|  |  | IO | 2 | SPARE
|  |  | IO | 3 | SPARE
Temp Sens | 4 | RTD-x |  | Ch# | EPICS Channel Name
Temp-A | 2 | RTD0 | READ | 0 | A1:ISI-T240_A_POD_TEMP_C | 11-blk1, 12-BRN, 13-RED | BLK->pin1, BRN->pin2, RED->pin3
|  |  | READ | 1 | A1:ISI-T240_B_POD_TEMP_C | 14-ORG, 15-YEL, 16-GRN | ORG->pin 7, YEL->pin 8, GRN->pin 9
Temp-B | 2 |  | READ | 2 | A1:PEM-LABORATORY_TEMP_C | 21-WHT, 22-GRY, 23-PUR | WHT->pin 1, GRY->pin 2, PUR->pin 3
|  |  | READ | 3 | A1:PEM-FRAME_TEMP_C | 24-BLU, 25-GRN, 26-YEL | BLU->pin 7, GRN->pin 8, YEL->pin 9
Temp-C | 2 |  | READ | 4 | A1:PEM-TANK_BASE_TEMP_C | 41-ORG, 42-RED, 43-BRN | ORG->pin 1, RED->pin 2, BRN->pin 3
|  |  | READ | 5 | A1:PEM-TANK_LID_TEMP_C | 44-BLK, 45-WHT, 46-GRY | BLK->pin 7, WHT->pin 8, GRY->pin 9

## Sheet1
|  | T0 | 100
|  | a | 0.0039083 |  |  | ref current | 0.001
|  | b | -5.775e-07
|  |  |  | T | R(T) | degC | V(1mA) | ADC (0.5V-16Bit)
243 | 100.2886 | 0.73841 |  | 24.5 | 109.5407 | 24.5 | 0.10954 | 7010.6029

## Versioning
v1 | initial release
v2 | LabView VI channel listing
v3 | initial ACORMAG release
v4 | add initial ISC channels
v5 | updated the ISC channels and modules accoridng to the .db files